# Reviewer search checklist

DEVICE SOFTWARE, CYBERSECURITY AND INTERFACES

Use the topic checks while searching the submission. Each checkbox identifies a specific check; Where and Search guide navigation, Why gives its purpose, and Details to check lists the information to look for.

Tick after examining the applicable check. A tick means examined, with acceptability judged separately. Use existing review notes for findings.

☐ U01 Confirm the evidence identity Where: Document cover; headers; configuration tables · Search: revision, build, configuration Why: Keeps evidence tied to the correct object. Details to check: Document/revision and original page/table; software/build; relevant time, population and configuration; distinguish printed and PDF page numbers.

☐ U02 Establish which expectation applies Where: Review scope; applicable guidance/standard; protocol · Search: scope, level, class, criterion Why: A checklist prompt alone does not make information mandatory. Details to check: Relevant device/function; documentation route; exact edition/section; requirement versus recommendation or sponsor commitment; justified alternative or nonapplicability.

☐ U03 Read the original content around search hits Where: Body text; tables; footnotes; figures; appendices · Search: synonyms, appendix, referenced ID Why: Keywords and document titles are only navigation clues. Details to check: Actual content; units; shared headers; row exceptions; referenced evidence; unreadable/unexamined scope. A keyword miss does not prove absence.

☐ U04 Separate what is claimed from what is demonstrated Where: Plan; execution record; result; linked verification · Search: planned, completed, actual, pass Why: A statement of completion may omit the underlying observation. Details to check: Plan versus result; expected versus actual; reported outcome; conditions; accessible supporting evidence. Supplied information can document a failed result.

☐ U05 Check conflicting and adverse evidence Where: Detailed results; deviations; changes; open-item lists · Search: failed, deferred, excluded, superseded Why: A favorable conclusion may conceal a material qualification. Details to check: Compatible IDs/versions/populations; failures and exceptions; contradictory values; governing evidence or unresolved conflict; basis for any bounded absence observation.

6.0.0 / search checklist 2.0 · 2026-10-06

# Topic finder — T00 to T15

Select a topic to jump to its checks. Page numbers refer to the full PDF; Word pagination may differ.

| Topic | What to search for | PDF pages |
| --- | --- | --- |
| T00 | Submission/module context and scope | 4 |
| T01 | Documentation level and software safety classification | 5 |
| T02 | Software purpose, functions, inputs and outputs | 6 |
| T03 | Software versions and configurations | 7 |
| T04 | Safety risk management | 8–9 |
| T05 | Software requirements | 10 |
| T06 | System and software architecture | 11 |
| T07 | Software design | 12 |
| T08 | Development, configuration management and maintenance practices | 13–14 |
| T09 | Software verification and validation | 15–16 |
| T10 | Version history and changes | 17 |
| T11 | Unresolved software anomalies | 18 |
| T12 | Third-party / SOUP / off-the-shelf software | 19 |
| T13 | Cybersecurity risk management process and report | 20 |
| T14 | Threat modeling | 21 |
| T15 | Cybersecurity risk assessment | 22 |

For clinical/analytical performance, manufacturing, biocompatibility or other domains, use the appropriate review resources. These 31 topics cover software and its interfaces.

# Topic finder — T16 to T30

Select a topic to jump to its checks. Page numbers refer to the full PDF; Word pagination may differ.

| Topic | What to search for | PDF pages |
| --- | --- | --- |
| T16 | Software bill of materials | 23 |
| T17 | Vulnerabilities and component support | 24 |
| T18 | Security implications of unresolved anomalies | 25 |
| T19 | Cybersecurity traceability | 26 |
| T20 | Cybersecurity measures and metrics | 27 |
| T21 | Security controls and requirements | 28–29 |
| T22 | Security architecture views | 30 |
| T23 | Cybersecurity testing | 31–32 |
| T24 | Cybersecurity labeling | 33 |
| T25 | Postmarket cybersecurity management | 34 |
| T26 | Cyber-device applicability statements | 35 |
| T27 | Usability and human factors | 36–37 |
| T28 | Standards and conformity claims | 38 |
| T29 | Software/device labeling | 39 |
| T30 | Machine-learning functions and controls | 40–41 |

For clinical/analytical performance, manufacturing, biocompatibility or other domains, use the appropriate review resources. These 31 topics cover software and its interfaces.

# T00 — Submission/module context and scope

Applies: Start here for every review; use modular-submission checks only when that route applies.

☐ T00.01 Identify the device and software functions under review Where: Cover letter; executive summary; device description · Search: applicant, device, module, intended use Why: Sets the boundary for every later check. Details to check: Applicant and device identity; software function or module; intended use; pathway; included and excluded functions.

☐ T00.02 Find the governing submission snapshot Where: Submission index; document register; amendment correspondence · Search: revision, amendment, supersedes, snapshot Why: Prevents combining superseded and current evidence. Details to check: Document titles and revisions; submitted dates; amendments; stated supersession; software/configuration addressed by each change.

☐ T00.03 Check that referenced evidence is accessible Where: Index; report cross-references; appendices and attachments · Search: appendix, attachment, report ID, see also Why: A listed document may supply no examinable content. Details to check: Referenced title/ID and revision; actual attachment; relevant pages, tables and images; missing or unreadable regions.

☐ T00.04 Identify earlier evidence being relied upon Where: Prior-submission references; bridging rationale; review correspondence · Search: prior submission, previously reviewed, unchanged Why: Earlier evidence may address a different device or scope. Details to check: Prior reference and decision scope; product/version; reused evidence; intervening changes; rationale linking it to this submission.

☐ T00.05 Locate unresolved commitments and later changes Where: Module correspondence; response letters; updated plans · Search: pending, will provide, follow-up, post-acceptance Why: Distinguishes completed evidence from future work. Details to check: Agreed deliverable and scope; target event/date; current status; submitted response; changed evidence or conditions after module acceptance.

☐ T00.06 Route material beyond software review Where: Contents; device description; performance and manufacturing sections · Search: clinical, analytical, manufacturing, electrical safety Why: Software coverage does not settle other device questions. Details to check: Relevant source and affected function; relationship to software evidence; appropriate review domain; any explicitly unexamined scope.

Related checks: T03, T09, T10, T28  |  Source pointers: S1 §III; S14 §VI

# T01 — Documentation level and software safety classification

Applies: Check the claimed documentation level and any claimed software safety class as separate decisions.

☐ T01.01 Find the documentation-level decision and rationale Where: Documentation-level evaluation; risk file · Search: Basic, Enhanced, documentation level Why: Determines the applicable software-documentation route. Details to check: Selected level; assessed device/functions; supporting failure and harm scenarios; cited guidance basis; rationale for the chosen route.

☐ T01.02 Check the harm assumptions behind the level Where: Hazard analysis; classification rationale · Search: before controls, serious injury, death, failure Why: A category label alone does not explain the decision. Details to check: Relevant failure modes; possible harm; conditions considered; treatment of risk controls; support for severity assumptions.

☐ T01.03 Find the software safety-class decision, if claimed Where: Safety-classification record; conformity documentation · Search: class A, class B, class C, 62304 Why: This classification scheme has its own scope and criteria. Details to check: Exact standard/edition; system or item classified; selected class; rationale; contributing software failures; external risk controls relied upon.

☐ T01.04 Check separately classified software items Where: Architecture; segregation rationale; class assignment · Search: segregation, software item, external control Why: A lower item class depends on the stated boundary. Details to check: Item boundaries; interfaces and dependencies; segregation mechanism; relied-on hardware or external controls; supporting evidence.

☐ T01.05 Compare classification with the actual device scope Where: Device description; safety and security risk files · Search: intended use, configuration, cybersecurity harm Why: Changes or omitted functions can undermine the rationale. Details to check: Same intended use and release; included functions; security-related harm; later changes; unresolved inconsistencies between classifications and risks.

Related checks: T04, T06, T13, T28  |  Source pointers: S1 §§V, VI.A; S10 §4.3

# T02 — Software purpose, functions, inputs and outputs

Applies: Apply to the software functions under review and to other functions that can affect them.

☐ T02.01 Identify intended use, users and workflow Where: Device description; intended-use statement; IFU · Search: intended user, population, workflow, automated Why: Gives requirements and tests their operating context. Details to check: Clinical purpose; intended users and population; environment; user actions; automated steps; software role in the device.

☐ T02.02 Check each material input and its limits Where: Interface specification; SRS; data dictionary · Search: input, format, units, range, validation Why: Input assumptions define meaningful test boundaries. Details to check: Input source; data type/format; units; allowed and validated ranges; quality checks; handling of missing, invalid or out-of-range data.

☐ T02.03 Check each material output and its recipient Where: Output specification; report examples; IFU · Search: output, result, display, recipient, units Why: An output can be technically correct yet misinterpreted downstream. Details to check: Meaning and units; format; receiving user/system; interpretation limits; error or invalid-result indications; links to underlying functions.

☐ T02.04 Identify algorithms and automated decisions Where: Software description; algorithm overview; workflow diagrams · Search: algorithm, calculation, decision, preprocessing Why: Locates functions needing requirements and performance support. Details to check: Method and purpose; processing sequence; assumptions; configurable settings; decision thresholds where used; relevant limitations.

☐ T02.05 Find supported platforms and deployment conditions Where: System requirements; installation guide; architecture · Search: hardware, operating system, hosting, cloud Why: A platform family may conceal untested configurations. Details to check: Hardware and OS versions; local/cloud deployment; required resources; dependencies; permitted configurations; tested versus supported limits.

☐ T02.06 Check external interfaces and other-function interactions Where: Interface-control documents; system diagrams; interoperability analysis · Search: protocol, counterpart, other function, data flow Why: Shared resources and interfaces can affect reviewed functions. Details to check: Source/target; protocol/version; data direction; counterpart products; timing/error behavior; shared resources; safety/effectiveness impact.

Related checks: T05, T06, T09, T27, T29  |  Source pointers: S1 §VI.B; S15 §§V–VI

# T03 — Software versions and configurations

Applies: Apply to each proposed, tested, released or historical baseline used in the evidence.

☐ T03.01 Identify the proposed release Where: Release manifest; version history; submission description · Search: proposed version, release, build ID, hash Why: Establishes the software object the review concerns. Details to check: Version and naming scheme; build ID/hash if used; build/release date; device/module identity; stated scope of a version-only identifier.

☐ T03.02 Identify the versions actually tested Where: Test-report headers; execution logs; configuration tables · Search: tested build, baseline, configuration Why: Results support the tested configuration first. Details to check: Software/build; hardware and OS; component versions; configuration settings; model/data version if relevant; report-to-run consistency.

☐ T03.03 Find the controlled configuration baseline Where: Configuration register; build manifest; toolchain record · Search: configuration item, compiler, dependency, baseline Why: A shared product version may hide different contents. Details to check: Configuration items and versions; build tools/settings where material; dependency baseline; reproducibility information; approved variants.

☐ T03.04 Check user-visible version identification Where: About screen; labels; IFU; service instructions · Search: version display, About, software identifier Why: Users and service staff need to identify the relevant release. Details to check: Displayed software version; access method; relationship to build identity; device identifiers kept distinct from software version.

☐ T03.05 Check changes between tested and proposed versions Where: Change log; impact assessment; regression report · Search: change since, freeze, bridge, equivalence Why: A different build needs a supported relevance argument. Details to check: Exact differences; affected functions and dependencies; risk impact; bridging rationale; repeated tests; dates as context, not proof of content identity.

Related checks: T09, T10, T16, T23, T29  |  Source pointers: S1 §§VI.H–I; S10 §§5.8, 8.1–8.3

# T04 — Safety risk management (1/2)

Applies: Apply to software-related risks within the device risk-management scope.

☐ T04.01 Find the risk plan and decision criteria Where: Risk-management plan; approved risk procedure · Search: acceptability, responsibility, overall residual Why: Risk judgments need an established method and scope. Details to check: Device/lifecycle scope; responsibilities; approval; risk-acceptability criteria; overall-risk method; verification and post-production arrangements.

☐ T04.02 Follow failure scenarios through to harm Where: Hazard analysis; use-related risk analysis; software risk analysis · Search: hazard, hazardous situation, sequence, harm Why: A failure label alone does not describe patient or user risk. Details to check: Intended use/misuse; causes and event sequence; hazardous situation; resulting harm; software contribution; initial severity/probability and their basis.

☐ T04.03 Check the selected risk controls Where: Risk-control table; requirements; design records · Search: control option, inherent safety, protective measure Why: The control should address the stated risk mechanism. Details to check: Control ID and rationale; affected hazard; implementation requirement/design; alternative options considered; reliance on warnings or external actors.

☐ T04.04 Find evidence that each relied-on control is implemented Where: Trace matrix; design verification; test records · Search: implemented, control verification, requirement ID Why: A proposed control may never have reached the tested release. Details to check: Control-to-requirement/design link; implemented version; verification method; executed result; unresolved implementation failures.

☐ T04.05 Find evidence that each relied-on control is effective Where: Effectiveness tests; validation; human-factors evidence · Search: effectiveness, residual risk, user task Why: Existence of a control does not establish risk reduction. Details to check: Risk-reduction objective; conditions and criterion; actual result; intended users/environment for human controls; remaining limitations.

☐ T04.06 Check risks introduced or changed by controls Where: Updated hazard analysis; change-impact review · Search: new risk, increased risk, side effect Why: A control can create another hazardous situation. Details to check: New hazards; changed likelihood/severity; affected functions; further controls; associated verification and residual-risk decisions.

Related checks: T05, T09, T11, T15, T27, T29  |  Source pointers: S1 §VI.C; S11 §§4.4–10; S10 §7

# T04 — Safety risk management (2/2)

Applies: Apply to software-related risks within the device risk-management scope.

☐ T04.07 Check individual residual-risk decisions Where: Risk evaluation; benefit-risk analysis where invoked · Search: residual, unacceptable, benefit-risk Why: Individual risk acceptance needs its own supporting argument. Details to check: Post-control risk and criterion; supporting evidence; justification where benefits are invoked; unresolved anomalies; necessary disclosures.

☐ T04.08 Find the overall risk conclusion and ongoing review Where: Risk-management report; production/post-production plan · Search: overall residual, release review, monitoring Why: Separate acceptable rows do not by themselves establish overall acceptability. Details to check: Combined residual-risk rationale; completed controls; review/approval; information sources; reassessment triggers; user disclosures; follow-up responsibilities.

Related checks: T05, T09, T11, T15, T27, T29  |  Source pointers: S1 §VI.C; S11 §§4.4–10; S10 §7

# T05 — Software requirements

Applies: Apply to the current requirements for the functions and configurations being reviewed.

☐ T05.01 Identify the governing requirements baseline Where: SRS cover/revision history; requirements export · Search: SRS, approved, baseline, revision Why: Trace links are meaningful only for the relevant requirements. Details to check: Document set and revision; approval; system/item scope; export date; changed or superseded requirements; consistent IDs and namespaces.

☐ T05.02 Check that functional requirements are specific Where: SRS; user-needs allocation; interface requirements · Search: shall, input, output, behavior, condition Why: Vague behavior cannot support a clear verification comparison. Details to check: Required behavior; triggering conditions; input/output; units and limits; boundary cases; error handling; applicable configurations.

☐ T05.03 Find measurable performance and constraint criteria Where: SRS performance sections; timing/resource specifications · Search: latency, accuracy, capacity, limit, acceptance Why: A test pass needs a defined target and conditions. Details to check: Threshold/range; units; load and operating conditions; resource constraints; rationale or higher-level need supporting material limits.

☐ T05.04 Locate safety and security requirements Where: Risk-control allocation; security requirements; SRS · Search: risk control, safety requirement, security requirement Why: Risk controls can disappear between analysis and implementation. Details to check: Requirement ID and text; linked hazard/threat/control; safety/security significance; responsible software item; associated verification method.

☐ T05.05 Follow requirements into design and executed verification Where: Trace matrix; SDS; test cases and results · Search: traceability, implements, verifies, test ID Why: A case-definition link does not establish an executed outcome. Details to check: Requirement/design/case/run IDs; compatible versions; actual outcome; uncovered requirements; design or tests lacking a justified requirement link.

☐ T05.06 Check revised requirements against earlier testing Where: Change records; SRS history; regression analysis · Search: revised requirement, impact, retest Why: An older result may not address a changed target. Details to check: Changed wording/criterion; effective version/date; affected controls and tests; rationale for reuse or retest; resulting evidence.

Related checks: T04, T07, T09, T19, T21  |  Source pointers: S1 §VI.D; S10 §5.2

# T06 — System and software architecture

Applies: Apply to the system and software boundaries supporting the reviewed functions.

☐ T06.01 Identify the system elements and allocated functions Where: Architecture diagram; architecture narrative · Search: module, item, layer, allocation Why: Shows where reviewed behavior is implemented. Details to check: Named elements; responsibilities; hardware/software allocation; third-party components; correspondence with the software description.

☐ T06.02 Trace internal interfaces Where: Interface-control documents; data-flow diagrams · Search: source, target, interface, data type Why: Hidden connections can invalidate isolation or test coverage. Details to check: Connected elements; direction; data/control flow; format and timing; error paths; consistency between diagrams and explanatory text.

☐ T06.03 Locate external systems and deployment boundaries Where: Deployment diagram; network architecture; service description · Search: cloud, hospital, external, hosted, service Why: The operating system can extend beyond the device enclosure. Details to check: External services and counterpart products; ownership; interface purpose; deployment configuration; included and excluded systems.

☐ T06.04 Check claimed segregation or isolation Where: Segregation rationale; architecture verification · Search: isolation, partition, separation, independence Why: A diagram boundary alone does not demonstrate separation. Details to check: Mechanism enforcing separation; shared resources; failure propagation; dependencies; supporting verification; links to classification and risk controls.

☐ T06.05 Reconcile architecture with requirements and the release Where: Configuration manifest; SRS; alternate architecture views · Search: component name, version, optional, other function Why: Different names or versions can hide a scope mismatch. Details to check: Element identity across views; proposed configuration; optional/other-function interactions; omitted elements; change impact and unresolved inconsistencies.

Related checks: T02, T04, T07, T14, T22  |  Source pointers: S1 §VI.E and Appendix B; S10 §5.3

# T07 — Software design

Applies: Use the established documentation route: detailed design is not a routine Basic-level submission request; examine it when supplied, applicable or needed for a specific question.

☐ T07.01 Locate the design for each material function Where: SDS; algorithm specification; detailed-design annex · Search: design element, implements, algorithm Why: Connects required behavior to the implementation approach. Details to check: Design element ID; software item; implemented requirement; logic/algorithm; assumptions; configuration scope.

☐ T07.02 Check interface and data-handling design Where: SDS interface sections; data structures; state diagrams · Search: state, transition, data structure, timeout Why: High-level architecture may omit decisive failure behavior. Details to check: Input/output structures; interface behavior; state transitions; error/timeout handling; exceptional conditions; safety/security implications.

☐ T07.03 Check that design links cover the relevant requirements Where: Requirement-to-design matrix; design review · Search: trace, unimplemented, orphan, allocation Why: A design description may omit or add behavior. Details to check: Requirement and design IDs; compatible revisions; uncovered requirements; untraced design elements; rationale for introduced behavior.

☐ T07.04 Find design-verification evidence Where: Design reviews; detailed-design verification records · Search: design verification, review finding, resolution Why: A design statement alone does not show it was checked. Details to check: Reviewed revision; method and criteria; findings; corrections; unresolved items; relationship to implementation and testing.

☐ T07.05 Check design changes against the tested baseline Where: SDS history; change-impact records; test configuration · Search: revision, approval, design change Why: Later design content may not be represented by earlier tests. Details to check: Initial and revised design scope; affected items/interfaces; governing baseline; impact assessment; relevant verification or retesting.

Related checks: T05, T06, T09, T21  |  Source pointers: S1 §VI.F; S10 §5.4; S13 §5.4

# T08 — Development, configuration management and maintenance practices (1/2)

Applies: Follow the applicable documentation level and conformity route. Internal process records are not automatic submission deliverables.

☐ T08.01 Find the claimed lifecycle and documentation route Where: Development plan; process description; conformity declaration · Search: lifecycle, process, 62304, declaration Why: Determines which process evidence the sponsor relies upon. Details to check: Lifecycle activities; responsibilities; deliverables; claimed clauses/edition; submitted procedures or descriptions; justified tailoring.

☐ T08.02 Check configuration and change control Where: Configuration-management plan; change procedure · Search: baseline, change control, status accounting Why: Uncontrolled changes undermine evidence identity. Details to check: Configuration identification; approval of changes; version/status tracking; source custody; release baseline; links to impact assessment and verification.

☐ T08.03 Check traceability practices Where: Development plan; traceability procedure; tool exports · Search: traceability, user need, requirement, test Why: Process links should support product-specific evidence chains. Details to check: Objects linked; forward/reverse scope; maintenance responsibility; change handling; tool/export configuration; actual release-specific application where relevant.

☐ T08.04 Find coding and development-tool controls Where: Coding standards; tool register; build procedure · Search: coding standard, compiler, static analysis, tool version Why: Development settings can affect repeatability and defects. Details to check: Applicable coding rules; tools/versions/settings; review or verification approach; deviations; external-component configuration controls.

☐ T08.05 Check maintenance and regression arrangements Where: Maintenance plan; regression policy; update process · Search: maintenance, impact analysis, regression, update Why: The release needs a controlled path for future changes. Details to check: Change/problem intake; impact analysis; regression selection; update verification; release approvals; supplier changes; responsibilities and commitments.

Related checks: T03, T09, T10, T12, T25  |  Source pointers: S1 §VI.G; S10 §§4.4, 5.1, 6, 8–9; S13 §§4–6, 8–9

# T08 — Development, configuration management and maintenance practices (2/2)

Applies: Follow the applicable documentation level and conformity route. Internal process records are not automatic submission deliverables.

☐ T08.06 Check problem resolution through closure Where: Problem-report procedure; linked change/test records · Search: problem report, investigate, resolution, trend Why: Closing a ticket does not demonstrate resolution of its effects. Details to check: Investigation; cause and impact; affected products/versions; corrective change; verification; communication; retained unresolved items.

☐ T08.07 Check legacy-software justification when invoked Where: Legacy assessment; gap analysis; remediation plan · Search: legacy, gap, continued use, remediation Why: A legacy label does not establish adequate supporting evidence. Details to check: Legacy boundary; risk evaluation; available evidence; process/evidence gaps; additional work; closure status; rationale for continued use.

Related checks: T03, T09, T10, T12, T25  |  Source pointers: S1 §VI.G; S10 §§4.4, 5.1, 6, 8–9; S13 §§4–6, 8–9

# T09 — Software verification and validation (1/2)

Applies: Apply the established documentation route. S1 §VI.H covers system protocols/results at Basic level and additionally unit/integration protocols/reports at Enhanced level; use targeted detail for the review question.

☐ T09.01 Find the verification and validation objectives Where: V&V plan; validation protocol; requirements matrix · Search: objective, user need, unit, integration, system Why: Establishes what each activity was intended to demonstrate. Details to check: Verification versus validation purpose; levels; user needs/requirements/controls covered; planned methods; scope exclusions and rationale.

☐ T09.02 Find the governing protocol and acceptance criteria Where: Approved protocol; test-case definitions; revision history · Search: expected result, criterion, approved, revision Why: An outcome needs a defined target and governing method. Details to check: Case ID; protocol revision; criterion and units; operating conditions; approval/execution dates; deviations or amended criteria.

☐ T09.03 Match the tested configuration to the proposed release Where: Test headers; execution logs; release manifest · Search: tested version, build, baseline, configuration Why: Evidence from a different configuration needs a bridge. Details to check: Software/build; platform/dependencies; tools and data; settings; proposed release; intervening differences; equivalence or retest evidence.

☐ T09.04 Find actual observations behind reported outcomes Where: Result tables; logs; appendices; original reports · Search: actual result, observed, measured, pass, fail Why: A Pass label may not show what was observed. Details to check: Case/run ID; actual value or behavior; units; criterion; reported outcome; source row/log; qualifications beside the result.

☐ T09.05 Check test conditions and boundary coverage Where: Methods; test-data description; table notes · Search: boundary, load, logging disabled, worst case Why: A favorable result may depend on restricted conditions. Details to check: Input ranges; load/configuration; relevant boundary/error conditions; footnotes; exclusions; justification for representativeness.

☐ T09.06 Reconcile execution counts and coverage Where: Execution register; result tables; coverage report · Search: planned, executed, blocked, deferred, retest Why: Overlap and different counting units can mislead. Details to check: Cases versus runs; source population; executed/pass/fail/blocked/deferred totals; repeated runs; overlap; unexplained differences.

Related checks: T03, T04, T05, T10, T11, T23  |  Source pointers: S1 §VI.H; S10 §§5.5–5.8

# T09 — Software verification and validation (2/2)

Applies: Apply the established documentation route. S1 §VI.H covers system protocols/results at Basic level and additionally unit/integration protocols/reports at Enhanced level; use targeted detail for the review question.

☐ T09.07 Follow failed and blocked tests to disposition Where: Failure log; deviation register; anomaly list · Search: failed, blocked, deviation, anomaly ID Why: Aggregate pass statements can conceal unresolved findings. Details to check: Failed criterion; actual observation; cause/impact; affected version; linked anomaly; disposition; release implications.

☐ T09.08 Check that claimed fixes have relevant retests Where: Retest logs; fix records; test summary · Search: fixed, rerun, retest, resolved Why: A fix claim does not supply verification of the fix. Details to check: Original failure; corrective change; fixed build; retest method/result; affected coverage; remaining failures or accepted limitations.

☐ T09.09 Check the regression-selection rationale Where: Change-impact assessment; regression plan/report · Search: selected tests, regression, impact, excluded Why: Repeating some tests may leave changed behavior uncovered. Details to check: Changed functions/interfaces/dependencies; affected risks/requirements; tests selected and omitted; rationale; executed outcomes.

☐ T09.10 Check whether the release conclusion matches the evidence Where: V&V conclusion; release review; unresolved-item register · Search: ready for release, conditional, deferred, limitation Why: A favorable conclusion may depend on unfinished work. Details to check: Supported version/scope; unresolved failures/anomalies; exclusions; residual conditions; referenced evidence actually available; future commitments.

Related checks: T03, T04, T05, T10, T11, T23  |  Source pointers: S1 §VI.H; S10 §§5.5–5.8

# T10 — Version history and changes

Applies: Apply when earlier versions, prior authorizations or changed configurations support the proposed release.

☐ T10.01 Find the version sequence and associated evidence Where: Version history; release notes; test/study reports · Search: version history, clinical version, last tested Why: Identifies which version each result supports. Details to check: Version/build sequence; dates; material changes; test, bench and clinical evidence associated with each baseline.

☐ T10.02 Identify differences from the last tested version Where: Change log; configuration comparison; release manifest · Search: change ID, from version, to version Why: A release label can conceal untested modifications. Details to check: Exact changed items/functions; dependencies and settings; previous/new versions; implemented versus planned changes.

☐ T10.03 Check safety, security and performance impact Where: Change-impact analysis; revised risk file · Search: no impact, affected control, new risk Why: A no-impact assertion needs a supporting scope and rationale. Details to check: Affected requirements/interfaces/controls; new or changed risks; downstream effects; basis for no-impact conclusions.

☐ T10.04 Find evidence supporting reuse of earlier results Where: Bridging rationale; equivalence analysis; regression report · Search: equivalent, bridge, unchanged, retest Why: Earlier results need a defensible link to current behavior. Details to check: Differences evaluated; unchanged evidence scope; technical rationale; repeated activities; actual results; residual limitations.

☐ T10.05 Check previously authorized and study versions separately Where: Prior-submission references; study configuration; current labeling · Search: authorized version, study build, marketed Why: Authorization and performance evidence may concern different baselines. Details to check: Prior decision scope; study/software configuration; candidate release; changed intended functions; bridge to each relied-on evidence source.

Related checks: T03, T04, T09, T16, T23  |  Source pointers: S1 §VI.I; S10 §§6.2–6.3, 7.4, 8.2

# T11 — Unresolved software anomalies

Applies: Apply to unresolved anomalies relevant to the proposed release and its evidence snapshot.

☐ T11.01 Identify the unresolved-anomaly population Where: Anomaly list; defect-register export; release review · Search: unresolved, open, query date, severity count Why: A few mentions do not establish the complete supplied list. Details to check: Software/build; list date; query filters; included statuses/severities; reported counts; received entries; scope exclusions.

☐ T11.02 Check the description and affected scope of each material anomaly Where: Individual defect records; reproduction notes · Search: anomaly ID, observed behavior, affected version Why: A short label may hide the relevant failure mechanism. Details to check: Unique ID; expected/actual behavior; trigger; discovery context; affected versions/configurations; investigation/root cause where available.

☐ T11.03 Check safety and effectiveness impact Where: Anomaly assessment; hazard analysis; release rationale · Search: impact, severity, hazard, unacceptable Why: Low defect priority does not by itself establish low device risk. Details to check: Severity taxonomy; patient/user consequences; linked risks; affected function; evidence and assumptions supporting the assessment.

☐ T11.04 Check security implications separately Where: Security anomaly assessment; vulnerability register · Search: security impact, weakness, vulnerability Why: A functional defect can also expose an attack path. Details to check: Security review criterion; relevant weakness; exploitability/impact rationale; linked vulnerability/control; unresolved consequences; see T18.

☐ T11.05 Check disposition, workaround and communication Where: Release review; anomaly disposition; IFU/advisory · Search: accepted, fixed, deferred, workaround, disclosed Why: Different disposition labels have different release implications. Details to check: Current status; effective version/date; justification; planned correction; practical workaround; user communication; remaining risk.

☐ T11.06 Reconcile anomalies with test failures and release statements Where: V&V failures; anomaly list; version history · Search: deferred test, failure ID, closed, release Why: The same issue may have inconsistent states across documents. Details to check: Matching IDs; compatible snapshots/configurations; failed/deferred tests; fixes/retests; related entries; explanation of apparent omissions or status conflicts.

Related checks: T04, T09, T10, T18, T29  |  Source pointers: S1 §VI.J; S10 §§5.8.2–5.8.3, 9

# T12 — Third-party / SOUP / off-the-shelf software

Applies: Apply to externally developed software relied upon by the device, using the relevant OTS/SOUP evidence route.

☐ T12.01 Identify each relevant external component and role Where: OTS/SOUP assessment; component register; SBOM · Search: supplier, component, version, patch, OTS, SOUP Why: Suitability evidence must concern the installed component. Details to check: Name; supplier; exact version/patch; function; device/software item; risk-control role; dependencies; classification used by the sponsor.

☐ T12.02 Check suitability and known limitations Where: Component assessment; supplier specifications · Search: suitability, limitation, restriction, intended use Why: General supplier capability may not support device-specific use. Details to check: Required behavior; operating limits; design limitations; known restrictions; intended device use; rationale and supporting evidence.

☐ T12.03 Check required environment and permitted configurations Where: Installation guide; compatibility matrix; configuration controls · Search: supported OS, permitted version, user configurable Why: User changes can invalidate assessed component combinations. Details to check: Hardware/OS requirements; compatible dependencies; permitted versions; configurable options; installation safeguards; prevention or handling of unsupported combinations.

☐ T12.04 Find the version-specific known-anomaly review Where: Supplier anomaly lists; issue tracker review; OTS assessment · Search: known bug, anomaly source, review date Why: A general component review can miss version-relevant problems. Details to check: Sources and query/review dates; component version; relevant anomalies; device impact; linked hazards/controls; disposition.

☐ T12.05 Check device-level verification of relied-on behavior Where: Integration tests; component tests; supplier reports · Search: integration, verified, supplier test, limitation Why: Supplier testing may cover a different use or environment. Details to check: Required behavior and risk controls; actual device configuration; test methods/results; limits of supplier evidence; remaining integration gaps.

☐ T12.06 Check support and replacement arrangements Where: Supplier lifecycle notice; maintenance plan; replacement assessment · Search: end of support, patch, replacement, continuity Why: Device support can outlast component support. Details to check: Component support date; device support horizon; monitoring/update responsibility; replacement or other mitigation; verification of changes; communicated constraints.

Related checks: T03, T04, T09, T16, T17  |  Source pointers: S2 §III; S10 §§5.3.3–5.3.4, 7.1.3

# T13 — Cybersecurity risk management process and report

Applies: Apply to the security-risk process for the device and the connected system within its defined scope.

☐ T13.01 Find the security-risk plan and report scope Where: Cybersecurity plan/report; system description · Search: scope, covered system, version, approval Why: A conclusion may exclude a relevant system element. Details to check: Document revisions; approval; device/system configuration; included assets/services; exclusions; lifecycle phases; relationship to the proposed release.

☐ T13.02 Check the security-risk method and decision criteria Where: Risk methodology; scoring procedure; assessment rubric · Search: acceptability, method, likelihood, residual Why: Scores and conclusions depend on their underlying method. Details to check: Evaluation method; criteria; assumptions; treatment of exploitability; risk acceptance roles; distinction between attack likelihood and patient-harm probability.

☐ T13.03 Check ownership and supplier responsibilities Where: Responsibility matrix; supplier agreement; security plan · Search: responsible, qualification, supplier, agreement Why: Allocating a task does not demonstrate it was completed. Details to check: Responsible roles and expertise; manufacturer/supplier boundaries; deliverables; escalation; evidence of completed responsibilities where claimed.

☐ T13.04 Find the interface with safety risk management Where: Security-to-safety mapping; safety risk file · Search: harm, safety transfer, patient impact Why: Technical security consequences can become device hazards. Details to check: Security event and affected function; hazardous situation/harm; linked safety-risk ID; transferred assumptions; corresponding controls and decisions.

☐ T13.05 Check the basis of the overall security conclusion Where: Security-risk report; open-findings register; release review · Search: residual risk, acceptable, unresolved, excluded Why: An overall statement can obscure accepted or unassessed risks. Details to check: Supporting assessments/tests; open findings; exclusions; external controls; accepted conditions; relevant release; support/end-of-service assumptions.

☐ T13.06 Check software support categories when IEC 81001-5-1 is claimed Where: Component classification; supplier responsibilities; risk-transfer records · Search: required, supported, maintained, transitional Why: Category and scope determine the stated maintenance responsibilities. Details to check: Component category; mapping of alternative terminology; manufacturer/operator duties; category changes and transferred risks; distinction from transitional status of the whole product.

Related checks: T04, T14, T15, T23, T25  |  Source pointers: S3 §V.A; S13 §§4.1–4.3, 7; S17

# T14 — Threat modeling

Applies: Apply to the actual device system, including relevant external services and lifecycle states.

☐ T14.01 Identify the modeled system and method Where: Threat-model report; architecture diagrams · Search: threat model, method, configuration, revision Why: A generic threat list may not represent this device. Details to check: Model date/revision; method and rationale; device/build scope; assets and security properties; included components and external systems.

☐ T14.02 Check entry points and trust boundaries Where: Data-flow diagrams; interface inventory; threat model · Search: trust boundary, entry point, interface, data flow Why: Unmodeled connections can leave attack paths unexplored. Details to check: Interfaces/endpoints; protocols; data direction; privilege boundaries; external dependencies; correspondence with deployed architecture.

☐ T14.03 Follow threats through attack paths Where: Threat tables; attack trees; misuse scenarios · Search: threat ID, attack vector, target, precondition Why: Threat names alone do not explain how harm could arise. Details to check: Threat ID; attacker capability/access; prerequisites; attack sequence; target asset; resulting impact; linked risk.

☐ T14.04 Check environmental assumptions and exclusions Where: Threat-model assumptions; deployment guide; labeling · Search: trusted network, assumption, excluded, transfer Why: A control argument may rely on conditions outside the device. Details to check: Each material assumption; justification or enforcement; external responsible actor; communication; consequences if violated; excluded-system rationale.

☐ T14.05 Check coverage beyond normal operation Where: Lifecycle/state diagrams; update and service threat models · Search: startup, update, recovery, service, decommission Why: Threat coverage can change across lifecycle states. Details to check: Relevant states; dormant/service interfaces; changed privileges; update/recovery paths; state-specific threats; scope of any exclusions.

☐ T14.06 Follow threats to controls, tests or explicit disposition Where: Threat-control matrix; risk assessment; security test plan · Search: mitigated by, control ID, test ID, unresolved Why: A modeled threat can remain without supporting mitigation evidence. Details to check: Threat/control/requirement/test links; implemented version; executed results; residual threats; rationale for accepted or unmitigated paths.

Related checks: T06, T15, T19, T21, T22, T23  |  Source pointers: S3 §V.A.1 and Appendix 2; S13 §§5.3, 7.2

# T15 — Cybersecurity risk assessment

Applies: Apply to individual security risks and the residual conditions relied on for the release.

☐ T15.01 Identify the risk and affected device scope Where: Security risk register; vulnerability assessment · Search: risk ID, threat, weakness, affected component Why: A vulnerability score may concern a different configuration. Details to check: Risk/vulnerability ID; linked threat/weakness; component and version; exposed function; affected release population; device/system consequences.

☐ T15.02 Check the scoring basis and assumptions Where: Risk worksheet; CVSS/rubric documentation · Search: vector, version, pre-mitigation, threshold Why: Different rubrics or assumptions yield incomparable scores. Details to check: Scoring system/version; vector or factors; device-specific rubric; criteria; exploit prerequisites; initial and residual assumptions.

☐ T15.03 Check what the mitigation actually changes Where: Risk-control rationale; architecture; mitigation assessment · Search: compensating control, post-mitigation, attack path Why: A lower score should have an explained mechanism. Details to check: Control and affected attack path; conditions for risk reduction; implementation scope; remaining exposure; changed score factors.

☐ T15.04 Find effectiveness evidence for the mitigation Where: Security tests; configuration verification; retest report · Search: effectiveness, retest, verified, bypass Why: A revised score is not proof of an effective control. Details to check: Control-specific test; relevant build/conditions; actual result; attempts to bypass or defeat the control where relevant; unresolved limitations.

☐ T15.05 Check residual and newly introduced risks Where: Residual-risk review; updated threat model · Search: residual, new risk, accepted, release condition Why: Mitigation can leave or introduce material concerns. Details to check: Remaining risk; applicable criterion; acceptance rationale; new risks; effective version/date; linked evidence and unresolved findings.

☐ T15.06 Check transferred risk and external controls Where: Risk-transfer agreement; IFU; service/support plan · Search: user responsibility, external control, end of service Why: Risk remains dependent on another actor's actions. Details to check: Recipient; required action/capability; conditions; communicated instructions; feasibility/evidence; duration; residual responsibility at support end.

Related checks: T04, T13, T17, T21, T23, T24  |  Source pointers: S3 §V.A.2; S13 §7

# T16 — Software bill of materials

Applies: Apply to the software/build and system-element scope represented by the SBOM.

☐ T16.01 Identify the SBOM artifact and build scope Where: SBOM file; explanatory report; build manifest · Search: SBOM, SPDX, CycloneDX, build, generation Why: Generation date alone does not establish release identity. Details to check: Format/specification version; author; generation tool/date; described product/build; connection to the proposed configuration.

☐ T16.02 Check component identity fields Where: SBOM entries; component register · Search: supplier, name, version, identifier, hash Why: Ambiguous names prevent reliable component matching. Details to check: Component name and supplier; exact version/patch; identifiers/hashes when used; relationship to the installed package; unresolved identity gaps.

☐ T16.03 Check dependency and system coverage Where: SBOM dependency graph; architecture; manifests · Search: dependency, transitive, excluded, cloud Why: Top-level inventory may omit relevant nested or external software. Details to check: Direct/transitive dependencies; relationships; device/cloud/intermediary scope; commercial/open-source/off-the-shelf components; omissions and rationale.

☐ T16.04 Check support and lifecycle information Where: SBOM support fields; supplier notices; support annex · Search: supported, end of support, end of life Why: Inventory alone does not establish maintainability. Details to check: Support status; end dates; affected component versions; source/date for lifecycle information; replacement or maintenance arrangements.

☐ T16.05 Match SBOM entries to vulnerability review Where: Component-to-vulnerability mapping; scan scope · Search: CVE, component match, query date, unmapped Why: A scan can miss components absent from its matching scope. Details to check: Same component/version identity; databases and query dates; matching method; unassessed or unmatched components; known-vulnerability disposition; see T17.

Related checks: T03, T12, T17, T19, T24  |  Source pointers: S3 §V.A.4

# T17 — Vulnerabilities and component support

Applies: Apply to known-vulnerability and component-support assessments for the relevant release population.

☐ T17.01 Check the vulnerability-search scope and currency Where: Scan report; database query record; component review · Search: database, query date, KEV, scan scope Why: An empty search result is bounded by its date and coverage. Details to check: Tools/databases and versions; query date; component population; matching rules; exploited-vulnerability checks; exclusions or inaccessible sources.

☐ T17.02 Match each material finding to the installed component Where: CVE/component mapping; SBOM; build configuration · Search: CVE, affected version, package, installed Why: A named vulnerability may concern a different component variant. Details to check: Vulnerability ID; component/supplier/version; affected range; installed configuration; mapping basis; unresolved identity ambiguity.

☐ T17.03 Check device-specific exploitability and impact Where: Vulnerability assessment; architecture; threat model · Search: not exploitable, disabled, prerequisite, impact Why: Generic severity does not resolve device-specific exposure. Details to check: Reachable interface; enabled feature; attacker access and prerequisites; affected function; safety/security consequences; evidence supporting dismissal.

☐ T17.04 Check mitigations and residual disposition Where: Mitigation record; risk assessment; retest · Search: fixed, compensating control, accepted, deferred Why: Disposition depends on effective action and its conditions. Details to check: Mitigation; implemented release; verification/retest; external reliance; residual risk; rationale for acceptance/deferral; planned correction.

☐ T17.05 Identify affected releases and unsupported components Where: Fielded-version inventory; supplier notices; support plan · Search: affected releases, unsupported, end of life Why: A supported candidate may coexist with vulnerable fielded versions. Details to check: Device/component versions; support end dates; deployed population; update/replacement coverage; unresolved gaps; user communication.

Related checks: T12, T15, T16, T23, T25  |  Source pointers: S3 §§V.A.4, V.A.6; S13 §§6.2, 9

# T18 — Security implications of unresolved anomalies

Applies: Apply to unresolved software anomalies that could have security consequences.

☐ T18.01 Check that security review covers the relevant anomaly list Where: Anomaly export; security assessment · Search: anomaly ID, reviewed list, snapshot Why: A security conclusion may omit later or differently filtered defects. Details to check: List/version/date; inclusion criteria; reviewed IDs; unreviewed items; reconciliation with T11 and the release snapshot.

☐ T18.02 Find the security impact of each material anomaly Where: Security anomaly analysis; threat/risk records · Search: security impact, denial of service, weakness, CWE Why: A tolerable functional defect can enable deliberate misuse. Details to check: Trigger; attacker influence; affected security property; weakness classification where used; device/system consequences; linked threats/risks.

☐ T18.03 Check the rationale for vulnerability classification Where: Anomaly-to-vulnerability decision; assessment criteria · Search: vulnerability, no security impact, not exploitable Why: A classification label needs a device-specific basis. Details to check: Decision criteria; rationale; configuration assumptions; supporting evidence; relationship to known vulnerabilities and attack paths.

☐ T18.04 Follow security-relevant anomalies to controls and verification Where: Vulnerability register; change record; security retest · Search: mitigation, retest, residual, accepted Why: Safety disposition alone may leave security consequences unresolved. Details to check: Shared anomaly identity; security control/fix; effective release; verification; residual consequences; acceptance/deferral rationale and communication.

Related checks: T11, T15, T17, T23  |  Source pointers: S3 §V.A.5; S13 §§5.8.5, 9

# T19 — Cybersecurity traceability

Applies: Apply to source-supported links within the relevant security evidence snapshot.

☐ T19.01 Identify the traceability export and namespaces Where: Trace matrix; requirements-tool export · Search: export date, baseline, ID, revision Why: Identical-looking IDs can refer to different objects or versions. Details to check: Export date; system/build; object types; ID namespaces; document revisions; included and excluded populations.

☐ T19.02 Follow threats and risks to requirements and controls Where: Threat-risk-control matrix; security requirements · Search: mitigated by, requirement ID, control ID Why: A risk can lack an implemented treatment despite a populated table. Details to check: Actual source-authored links; threat/risk identity; control/requirement text; accepted exceptions; missing or contradictory targets.

☐ T19.03 Follow controls to architecture and affected components Where: Architecture mapping; component/SBOM links · Search: architecture element, component ID, implements Why: A generic control may not protect the relevant asset or interface. Details to check: Implementing element; protected asset/interface; component/version; configuration; matching names/IDs across views.

☐ T19.04 Follow controls to executed tests Where: Trace matrix; test definitions; run results · Search: test case, run ID, verified, result Why: A link to a case is not evidence of execution. Details to check: Control/requirement ID; test ID; execution record; tested build; actual outcome; unexecuted or failed tests.

☐ T19.05 Check reverse coverage and dangling links Where: Reverse trace export; linked source records · Search: untraced, orphan, missing target, bidirectional Why: Selected forward links cannot establish complete coverage. Details to check: Scope actually checked; controls/tests without a supported source; unresolved targets; references whose contents are unavailable; version mismatches.

Related checks: T05, T14, T15, T16, T21, T23  |  Source pointers: S3 §§V.A–V.C and Appendix 2; S13 §§5.2–5.7

# T20 — Cybersecurity measures and metrics

Applies: Apply where security-maintenance metrics are supplied or relevant to stated commitments; distinguish new-product data limits.

☐ T20.01 Find each metric's definition and population Where: Metrics report; dashboard specification · Search: definition, denominator, reporting period Why: A percentage is uninterpretable without its population. Details to check: Numerator/denominator; device/version population; reporting period; units; inclusions/exclusions; source of the figures.

☐ T20.02 Check the identification-to-patch interval Where: Vulnerability/patch register; metrics report · Search: identification, discovery, confirmation, patch release Why: Different starting events produce different timing claims. Details to check: Start/end event definitions; dates; unit; affected vulnerabilities; calculation method; exceptions; distinction from deployment time.

☐ T20.03 Check patch-to-deployment and field coverage Where: Deployment report; fielded-version inventory · Search: deployed, installed, fielded, percentage patched Why: Releasing a patch does not show that devices received it. Details to check: Patch-release versus installation dates; eligible/updated population; affected versions; deployment evidence; offline or excluded devices.

☐ T20.04 Check zero, unavailable and nonapplicable results Where: Metric notes; supporting register; plan · Search: zero, no data, unavailable, not applicable Why: These states imply different evidence and limitations. Details to check: Explicit value and reason; actual data availability; product/lifecycle context; calculation scope; planned measurement where relevant.

☐ T20.05 Compare metrics with management commitments Where: Security-management plan; update policy; metrics · Search: target, routine, exceptional, service level Why: Apparent performance can depend on incompatible clocks or scope. Details to check: Same trigger, interval and population; stated target; reported result; missed targets; explanations; version-specific risk context.

Related checks: T17, T23, T25  |  Source pointers: S3 §V.A.6

# T21 — Security controls and requirements (1/2)

Applies: Select controls relevant to the device assets, interfaces, threats and deployment conditions; do not assume every mechanism applies.

☐ T21.01 Check authentication of users, devices and services Where: Authentication design; interface specification; security tests · Search: authentication, identity, credential, default password Why: Each trusted entity needs a supported identity mechanism. Details to check: Entity/interface covered; mechanism; initial credentials; credential changes/revocation; default/hardcoded credentials; failed-authentication behavior; verification.

☐ T21.02 Check authorization and privilege boundaries Where: Role/permission matrix; access-control design; tests · Search: role, authorization, least privilege, service access Why: An authenticated entity may still have excessive access. Details to check: Roles; permitted actions/resources; privileged/service functions; enforcement points; privilege changes; denied-action tests; exceptional access.

☐ T21.03 Check cryptographic protection of relevant flows and assets Where: Crypto specification; security architecture; test report · Search: algorithm, protocol, mode, key length Why: An algorithm name alone leaves protection scope unclear. Details to check: Protected data/flow; algorithm/protocol/version; mode/settings; key strength; endpoints; negotiated or fallback behavior; implementation/test evidence.

☐ T21.04 Check key and certificate management Where: Key-management design; provisioning/update procedures · Search: key generation, storage, certificate, revocation Why: Strong algorithms depend on protected, managed secrets. Details to check: Key creation/distribution; storage and access; device uniqueness/shared-key scope; replacement/revocation; certificate validation; responsibility and verification.

☐ T21.05 Check integrity of code and data Where: Boot/update design; data-integrity requirements; tests · Search: signature, integrity, authenticity, tamper Why: Corrupted or substituted content can affect device behavior. Details to check: Protected code/data; origin and integrity checks; verification stage; rejection behavior; relevant version metadata; tamper/failure test results.

☐ T21.06 Check confidentiality of sensitive data Where: Data-flow inventory; storage/transport design · Search: confidential, at rest, in transit, privacy Why: Protection can stop at an intermediate system or storage location. Details to check: Sensitive assets; storage/transmission paths; authorized access; protection endpoints; intermediaries; retention/deletion where relevant; verification scope.

Related checks: T05, T14, T15, T19, T22, T23  |  Source pointers: S3 §V.B.1 and Appendix 1; S13 §§5.2–5.7

# T21 — Security controls and requirements (2/2)

Applies: Select controls relevant to the device assets, interfaces, threats and deployment conditions; do not assume every mechanism applies.

☐ T21.07 Check security event detection and logs Where: Logging specification; administrator guide; tests · Search: event, audit log, timestamp, retention Why: Users may rely on logs that omit decisive events. Details to check: Events captured; timestamp/time source; fields/location; access and tamper protection; retention; alerting/review actions; behavior if logging fails.

☐ T21.08 Check safe behavior during security failures Where: Resilience design; denial-of-service tests; hazard analysis · Search: safe state, degraded, availability, resource exhaustion Why: Security controls can impair essential device functions. Details to check: Relevant failure/attack conditions; maintained functions; degraded/safe state; resource limits; recovery criteria; safety links; actual tests.

☐ T21.09 Check backup and recovery Where: Backup/restore procedure; recovery tests · Search: backup, restore, recovery, configuration Why: Recovery must restore a usable and controlled state. Details to check: Data/configuration covered; protection; access; restore prerequisites; restored version/state; failed-restore handling; demonstrated recovery.

☐ T21.10 Check update authenticity and interruption handling Where: Update design; rollback/recovery tests · Search: signed update, rollback, interruption, partial install Why: An update may leave untrusted or inconsistent software. Details to check: Update source and authentication; compatibility/version checks; interrupted-install state; rollback policy; recovery; signing-key protection; actual failure-path tests.

Related checks: T05, T14, T15, T19, T22, T23  |  Source pointers: S3 §V.B.1 and Appendix 1; S13 §§5.2–5.7

# T22 — Security architecture views

Applies: Check architecture views appropriate to the device system and relevant security use cases.

☐ T22.01 Find the global system view Where: Security architecture; deployment diagrams · Search: global view, cloud, intermediary, network Why: Shows the complete operating system and its trust boundaries. Details to check: Device and external elements; owners; trust boundaries; assets; dependencies; diagram revision and configuration; consistency with T06.

☐ T22.02 Check potential effects across multiple devices or patients Where: Multi-patient view; fleet/cloud architecture; threat model · Search: multi-patient, shared service, fleet, propagation Why: A shared dependency can expand the impact of compromise. Details to check: Shared assets/credentials/services; paths between devices; scale of effect; isolation and containment controls; rationale if this view is inapplicable.

☐ T22.03 Follow update and patch delivery end to end Where: Update architecture; distribution/service diagrams · Search: update server, distribution, signing, patch path Why: Protection may fail outside the device's local updater. Details to check: Source through intermediaries to device; ownership/trust; payload and metadata; authentication; delivery states; verification and recovery points.

☐ T22.04 Check security use cases across operating states Where: Security use-case views; state diagrams · Search: programming, alarming, standby, service, recovery Why: Clinical and operational states can expose different paths. Details to check: Relevant state and clinical function; end-to-end communication; enabled/dormant interfaces; privileges; controls; safety consequences; exclusions.

☐ T22.05 Check the detail on each relevant communication path Where: Interface/port tables; diagram annotations · Search: endpoint, port, protocol, session, default Why: A line between boxes omits the facts needed to assess protection. Details to check: Endpoints; direction; payload; protocol/version/port; authentication; cryptography; session handling; defaults; linked control/test/component IDs.

Related checks: T06, T14, T19, T21, T23, T24  |  Source pointers: S3 §V.B.2 and Appendix 2

# T23 — Cybersecurity testing (1/2)

Applies: Select test activities against the applicable security requirements, threat model and device boundaries.

☐ T23.01 Identify test activities, objectives and boundaries Where: Security test plan; report scope; architecture · Search: scope, objective, exclusion, test type Why: Different methods address different security questions. Details to check: Requirements, threats and interfaces covered; excluded systems; rationale; test activity type; connection to the threat model.

☐ T23.02 Check the tested configuration and tool settings Where: Report header; environment section; tool output · Search: build, tool version, configuration, database date Why: Results depend on the release, environment and test settings. Details to check: Software/build; components; enabled interfaces; environment; tool versions/settings; vulnerability-source dates; test dates; release bridge.

☐ T23.03 Check security-requirement and threat-mitigation tests Where: Requirements tests; threat-driven tests; control matrix · Search: requirement ID, mitigation, boundary, effectiveness Why: Implementing a control does not demonstrate resistance to the modeled threat. Details to check: Requirement/control/threat IDs; method; conditions; boundary assumptions; criterion; actual result; failed or excluded cases.

☐ T23.04 Check vulnerability-testing methods and coverage Where: Scan/analysis reports; robustness/fuzz tests · Search: malformed input, fuzz, attack surface, static, dynamic Why: A single scan does not cover all relevant weaknesses. Details to check: Methods and rationale; attack-surface coverage; scanning, static/dynamic and binary composition analysis; malformed-input/fuzz testing; vulnerability chaining where relevant; target interfaces/components; tool scope; findings and exclusions.

☐ T23.05 Check penetration-test expertise, independence and effort Where: Original penetration-test report; tester qualifications · Search: penetration, independence, expertise, duration Why: The strength of an exercise depends on who tested what and how. Details to check: Tester relationship to developers; qualifications; duration/effort; access/knowledge assumptions; methods; tested scope; original report availability.

☐ T23.06 Find the actual findings and observations Where: Original test reports; findings register; evidence annexes · Search: finding ID, observation, severity, evidence Why: A favorable executive statement can omit actionable findings. Details to check: Finding IDs; evidence; severity method; affected asset/build; reproduction conditions; reported counts; linked risk assessment.

Related checks: T03, T14, T15, T17, T19, T21, T25  |  Source pointers: S3 §V.C; S13 §§5.7–5.8

# T23 — Cybersecurity testing (2/2)

Applies: Select test activities against the applicable security requirements, threat model and device boundaries.

☐ T23.07 Check fixes against relevant retest results Where: Remediation register; retest report; fixed-build record · Search: corrected, fixed build, retested, remaining Why: A closed finding may have no demonstrated correction. Details to check: Finding-to-fix link; changed build; retest scope/method; actual result; regression effects; findings not retested.

☐ T23.08 Check accepted, deferred and excluded findings Where: Residual-risk review; test conclusion; release plan · Search: accepted, deferred, excluded, post-release Why: Remaining findings can qualify the release conclusion. Details to check: Each material disposition; risk rationale; compensating controls; affected release; user actions; future commitments; timing and verification plan.

Related checks: T03, T14, T15, T17, T19, T21, T25  |  Source pointers: S3 §V.C; S13 §§5.7–5.8

# T24 — Cybersecurity labeling

Applies: Apply to security-related instructions and disclosures needed for the intended users, administrators and service personnel.

☐ T24.01 Check deployment prerequisites and recommended controls Where: IFU security section; administrator guide; supplied disclosure questionnaire · Search: secure network, firewall, endpoint, ports Why: Risk arguments may depend on the user's environment. Details to check: Required/recommended infrastructure; network diagram; ports/protocols/direction; approved endpoints; controls users must configure; responsibility boundaries.

☐ T24.02 Check version identification and SBOM access Where: IFU; version screen; SBOM delivery instructions · Search: version, build, SBOM, portal Why: Users need information for their actual device configuration. Details to check: How to identify software; matching SBOM version; delivery/access method; intended recipient; actual artifact versus a portal reference.

☐ T24.03 Check incident recognition and response instructions Where: Security manual; incident-response section · Search: security event, log, incident, contact Why: Detection features are useful only if users can act on them. Details to check: Event/alert meaning; log location/format; user actions; escalation/contact; evidence preservation; safe operation during an incident.

☐ T24.04 Check secure setup, backup and recovery instructions Where: Installation guide; backup/restore procedure · Search: secure default, backup, restore, update Why: Reliance on administrative action needs usable instructions. Details to check: Initial configuration; account/access setup; backup protection; recovery steps; update process; failure handling; role-specific responsibilities.

☐ T24.05 Check support dates and component limitations Where: IFU support section; lifecycle notices; supplier annex · Search: support end, component support, end of service Why: Device and component support promises may differ. Details to check: Device/version horizon; relevant component end dates; replacement/update coverage; support-end actions; consistency with SBOM and management plan.

☐ T24.06 Check risk transfers, unsafe changes and decommissioning Where: Warnings; security limitations; sanitization procedure · Search: risk transfer, configuration change, sanitize, retire Why: General warnings may omit actions needed for the risk argument. Details to check: Specific user controls; changes increasing risk; residual risks; service/retirement steps; data/credential sanitization; feasibility and human-factors links.

Related checks: T15, T16, T21, T22, T25, T27, T29  |  Source pointers: S3 §VI.A and Appendix 1; S13 §5.8.2 and Annex E

# T25 — Postmarket cybersecurity management

Applies: Apply to the security-management commitments for supported and fielded device versions.

☐ T25.01 Find ownership and vulnerability monitoring arrangements Where: Postmarket security plan; PSIRT procedure · Search: PSIRT, monitoring, source, frequency Why: Detection depends on assigned and sustained activity. Details to check: Responsible roles; monitoring sources/frequency; component/SBOM linkage; affected-product identification; escalation and resourcing.

☐ T25.02 Check vulnerability disclosure and intake Where: Disclosure policy; intake procedure; public contact · Search: coordinated disclosure, intake, advisory Why: External findings need a defined path into action. Details to check: Reporting channels; triage; communication responsibilities; coordination; advisory contents; tracking through disposition.

☐ T25.03 Check routine and exceptional patch commitments Where: Patch policy; management plan · Search: routine, out-of-cycle, trigger, timeframe Why: Different risk situations can require different update paths. Details to check: Trigger definitions; timelines and units; prioritization; verification; approval; exceptional handling; distinction between planned targets and completed action.

☐ T25.04 Check deployment to fielded versions Where: Update architecture; field-service plan; version inventory · Search: deployment, fielded, offline, supported Why: A verified patch may not reach affected devices. Details to check: Supported release population; delivery mechanism; installation responsibility; compatibility; deployment confirmation; offline/failed updates; communications.

☐ T25.05 Check support and end-of-service transitions Where: Support policy; end-of-service plan; IFU · Search: supported, maintained, end of service, transfer Why: Risk can remain after routine maintenance ends. Details to check: Product/component horizons; notification; affected fielded versions; replacement/retirement options; residual risks; transferred responsibilities.

☐ T25.06 Compare commitments with available execution evidence Where: Metrics; incident/change records; prior updates · Search: completed, deployed, target, overdue Why: A plan and a completed activity support different claims. Details to check: Relevant event/version; stated target; actual action/date; exceptions; unresolved commitments; standards/process claims and their stated scope.

Related checks: T17, T20, T23, T24, T26  |  Source pointers: S3 §VI.B; S13 §§4.1.7–4.1.8, 6, 9; S16 §§V–VI

# T26 — Cyber-device applicability statements

Applies: Check factual characteristics and the applicability rationale against the relevant legal criteria; an unchecked box is not a legal determination.

☐ T26.01 Find the stated cyber-device applicability decision Where: Applicability statement; submission cover; legal-basis section · Search: cyber device, 524B, applicability Why: Identifies the decision and evidence route being claimed. Details to check: Applicable submission/device scope; stated conclusion; cited legal basis; date; underlying rationale; unresolved qualifications.

☐ T26.02 Check the software characteristic relied upon Where: Device description; software inventory; architecture · Search: includes software, validated, installed, authorized Why: The factual software scope must match the classification rationale. Details to check: Included software; installation/validation/authorization relationship; device functionality; relevant components and boundaries.

☐ T26.03 Check internet-connectivity capability Where: Interface specification; network/service architecture · Search: internet, indirect, optional, disabled, service port Why: A marketing label may omit actual connectivity capability. Details to check: Direct/indirect paths; connectors and counterpart systems; optional/service interfaces; enabled/disabled configurations; evidence supporting the claimed capability.

☐ T26.04 Check vulnerable technological characteristics Where: Architecture; threat model; security assessment · Search: technological characteristic, vulnerability, susceptible Why: Applicability needs the actual device characteristics. Details to check: Relevant hardware/software features; exposure and vulnerability rationale; configuration assumptions; consistency with connectivity and software descriptions.

☐ T26.05 Locate the evidence relied on under the applicable route Where: Applicability cross-references; management plan; SBOM · Search: monitoring plan, update process, SBOM Why: A document reference alone supplies no examinable content. Details to check: Actual monitoring/disclosure plan; update/security processes; SBOM; relevant device/version; specific applicability basis for each requested element.

Related checks: T02, T06, T16, T24, T25  |  Source pointers: S3 §VII

# T27 — Usability and human factors (1/2)

Applies: Use the justified human-factors submission route; examine validation details when applicable to the device, changes and use-related risks.

☐ T27.01 Find the user, environment and interface scope Where: Use specification; device description; IFU · Search: user group, environment, UI, training Why: Evidence must represent the people and conditions of use. Details to check: Intended user groups; capabilities; environments; tasks; software/UI/labeling versions; training assumptions; known use problems.

☐ T27.02 Find the claimed submission category and rationale Where: HF category assessment; change analysis · Search: category, rationale, new device, modified device Why: The applicable evidence route depends on the device and changes. Details to check: Claimed category; device/change scope; use-related risk reasoning; cited criteria; supporting analysis; unresolved applicability.

☐ T27.03 Identify critical tasks and their risk links Where: Task analysis; use-related risk analysis · Search: critical task, use error, hazardous situation Why: A study may omit the task that creates the risk. Details to check: Task ID and sequence; foreseeable errors; resulting harm; critical-task rationale; linked controls; scenarios selected for evaluation.

☐ T27.04 Check formative findings and design responses Where: Formative studies; design-change records · Search: formative, observed problem, design change Why: Earlier use problems may remain in the final interface. Details to check: Users/tasks studied; findings; selected design controls; changes made; verification/evaluation; issues carried into validation.

☐ T27.05 Check participant and scenario representativeness Where: Validation protocol/report; recruitment criteria · Search: participant, group, scenario, sample rationale Why: Completion counts are meaningful only for a relevant study. Details to check: User groups and counts; selection rationale; task/scenario coverage; environments; actual versus planned conditions; exclusions and deviations.

☐ T27.06 Check training, assistance and final interface fidelity Where: Study methods; moderator script; test configuration · Search: training, decay, prompting, assistance, final UI Why: Assistance or unfamiliar configurations can distort observed performance. Details to check: Training type/timing; user access to instructions; moderator intervention; UI/labeling/software versions; differences from marketed use.

Related checks: T02, T04, T09, T24, T29, T30  |  Source pointers: S4 §§5–8; S5 §§IV–V; S12 §§5.1–5.9

# T27 — Usability and human factors (2/2)

Applies: Use the justified human-factors submission route; examine validation details when applicable to the device, changes and use-related risks.

☐ T27.07 Find observed errors, close calls and difficulties Where: Task-level results; observation logs; interviews · Search: use error, close call, difficulty, root cause Why: Successful completion can conceal unsafe interaction. Details to check: Task/user group; actual behavior; assistance; errors and near misses; causes; participant feedback; relationship to critical tasks.

☐ T27.08 Check residual use-risk reasoning and remaining changes Where: HF conclusion; risk file; revised labeling/UI · Search: residual risk, unresolved, redesign, additional validation Why: A favorable overall statement must address material findings. Details to check: Finding-specific risk analysis; further design changes; justification for remaining risk; supporting evidence; need for further evaluation after changes.

Related checks: T02, T04, T09, T24, T29, T30  |  Source pointers: S4 §§5–8; S5 §§IV–V; S12 §§5.1–5.9

# T28 — Standards and conformity claims

Applies: Apply to each actual conformity, partial-use or equivalent-evidence claim; verify recognition for the relevant review date.

☐ T28.01 Identify the exact claim and product scope Where: Declaration; standards list; submission statement · Search: declaration, conformity, general use, edition Why: Mentioning a standard does not define the claimed evidence route. Details to check: Exact standard/edition/amendment; full or selected-clause claim; device/software identity; scope/options/classes; declaration date and signatory if supplied.

☐ T28.02 Check recognition and transition details where relied upon Where: FDA recognition entry; supplemental information · Search: recognition number, extent, transition, edition Why: Recognition can differ by edition, clause and date. Details to check: Dated official entry; recognized edition; extent/exclusions; supplemental information; transition conditions; match to the sponsor's actual claim.

☐ T28.03 Check exclusions, deviations and alternatives Where: Clause matrix; declaration limitations; rationale · Search: excluded, deviation, not applicable, equivalent Why: A broad claim can conceal important limitations. Details to check: Specific clauses/options affected; reasons; device scope; alternative method/evidence; limits of the resulting conformity claim.

☐ T28.04 Locate the supporting evidence Where: Referenced reports; test records; conformity attachments · Search: report number, evidence, completed, pending Why: A declaration or citation does not supply all underlying facts. Details to check: Report identity/revision; relevant contents; methods/results; completion versus future testing; access limits; relation to claimed clauses.

☐ T28.05 Check tested-product and edition consistency Where: Original reports; declaration; release/change records · Search: tested version, edition, equivalent, bridge Why: Evidence may concern a different product or standard edition. Details to check: Tested device/software; report edition/options; marketed release; changes; equivalence/bridging rationale; unresolved mismatches.

☐ T28.06 Check laboratory claims where relied upon Where: Laboratory report; accreditation certificate/scope · Search: accredited, scope, laboratory, method Why: Accreditation claims are limited to a stated scope and time. Details to check: Laboratory identity; relevant accreditation scope; validity at testing; method covered; report authenticity/provenance; limitations of the claim.

Related checks: T01, T03, T09, T23, T27  |  Source pointers: S6 §§IV–VIII; S7

# T29 — Software/device labeling

Applies: Apply to labeling supporting the reviewed software functions; use T24 for detailed cybersecurity instructions.

☐ T29.01 Match labeling to the device and software release Where: IFU cover/history; version screen; release manifest · Search: IFU revision, software version, About Why: A current document revision can still describe an older release. Details to check: Label/IFU revision; software identified; intended device/configuration; user-visible version; controlled draft/final status; discrepancies.

☐ T29.02 Check operating instructions and limits Where: IFU; technical description; training materials · Search: operating range, limitation, input, output Why: User-facing claims should match requirements and validation. Details to check: Intended users/use; input/output interpretation; operating limits; warnings; supported functions/configurations; consistency with tested conditions.

☐ T29.03 Check permitted external components and configuration changes Where: Installation guide; compatibility list; IFU · Search: permitted version, compatible, update, warning Why: User-selected components can alter validated behavior. Details to check: Allowed component/OS versions; installation or update restrictions; compatibility checks; user warnings; handling of unsupported combinations.

☐ T29.04 Check anomaly and residual-risk disclosures Where: IFU warnings; anomaly notice; risk-management report · Search: known limitation, anomaly, residual risk Why: Accepted risk may depend on information reaching users. Details to check: Relevant anomaly/risk; affected release; consequence; action/workaround; disclosure scope; consistency with accepted risk controls.

☐ T29.05 Check instructions used as risk controls Where: Risk-control matrix; IFU; human-factors evidence · Search: information for safety, warning, critical task Why: Wording alone does not establish that the control works. Details to check: Exact user action; intended user/environment; clarity and location; linked risk; effectiveness evidence; network-failure/security instructions where relevant.

Related checks: T02, T03, T04, T11, T24, T27  |  Source pointers: S1 §§VI.B, VI.J; S4 §§5, 8; S11 §§7.3, 8

# T30 — Machine-learning functions and controls (1/2)

Applies: Apply to machine-learning functions. PCCP checks apply when a plan is proposed or relied on; S9 is draft background, not a finalized requirement.

☐ T30.01 Identify the model and processing pipeline Where: Model description; artifact manifest; software configuration · Search: model version, artifact hash, framework, locked Why: Performance belongs to a particular model and pipeline. Details to check: Model identity/version/hash where used; framework; preprocessing/postprocessing; decision settings; locked/adaptive behavior; proposed versus tested artifact.

☐ T30.02 Check data provenance and intended-population coverage Where: Dataset description; provenance records; site tables · Search: dataset version, site, dates, population Why: Validation may not represent intended use. Details to check: Data sources; collection sites/time; selection criteria; relevant populations/conditions; dataset version; exclusions; relationship to intended use.

☐ T30.03 Check training, tuning and test independence Where: Partition records; study design; data-flow description · Search: train, tune, test, overlap, external validation Why: A split label does not rule out information leakage. Details to check: Partition method; patient/site/time overlap; separation of development and testing; model-selection use; test reuse; external cohort identity where claimed.

☐ T30.04 Check reference labels and data handling Where: Annotation protocol; reference-method description; data-quality report · Search: reference standard, ground truth, annotation, missing Why: Performance depends on trustworthy targets and processing. Details to check: Reference method; annotator/adjudication process where relevant; label quality; missing/invalid data; preprocessing; exclusions and their effect on evaluation.

☐ T30.05 Find performance targets and actual results Where: Validation protocol; performance tables; statistical analysis · Search: metric, threshold, denominator, confidence interval Why: An aggregate favorable metric can lack a justified target. Details to check: Metric definition; acceptance criterion and rationale; tested cohort/model; sample counts; actual result; uncertainty; failed criteria; conditions.

☐ T30.06 Check subgroup, site and condition performance Where: Stratified results; external-validation report · Search: subgroup, site, stratified, outlier Why: Overall results can conceal important differences. Details to check: Relevant subgroup/site definitions; denominators; estimates and uncertainty; exclusions; underrepresented groups; limitations on conclusions.

Related checks: T02, T03, T04, T09, T10, T27, T28  |  Source pointers: S8 §§VI–VIII and Appendix A; S9 §§V–XII (draft background only)

# T30 — Machine-learning functions and controls (2/2)

Applies: Apply to machine-learning functions. PCCP checks apply when a plan is proposed or relied on; S9 is draft background, not a finalized requirement.

☐ T30.07 Check limitations and user interpretation Where: Model labeling; transparency material; HF evidence · Search: limitation, uncertainty, explanation, human oversight Why: Users need to understand the output's intended role and limits. Details to check: Supported inputs/populations; known failure conditions; output interpretation; user actions; disclosures; consistency with observed performance.

☐ T30.08 Check ML-specific safety and security controls Where: ML risk assessment; robustness tests; monitoring plan · Search: data integrity, model integrity, drift, failure mode Why: Data or model changes can alter device behavior. Details to check: Relevant hazards/threats; model/data integrity controls; robustness evidence; monitoring triggers where relied upon; response actions; residual risks.

☐ T30.09 Check proposed or authorized model changes Where: PCCP; modification protocol; impact assessment · Search: PCCP, modification, retraining, acceptance, rollback Why: A proposed plan does not establish authorization or completed validation. Details to check: Plan status; specified modifications; data/retraining controls; evaluation criteria; update procedures; impact assessment; modification-to-test links; version control and user communication.

Related checks: T02, T03, T04, T09, T10, T27, T28  |  Source pointers: S8 §§VI–VIII and Appendix A; S9 §§V–XII (draft background only)

# Source pointers — 1

Use the sections cited with each topic. These are reading pointers for authored checks, not a clause-by-clause compliance checklist. Confirm the applicable edition, recognition and evidence route for the case.

S1 — Content of Premarket Submissions for Device Software Functions
June 2023; final guidance  Official source
Local copy: fda/Premarket-Software-Functions-Guidance.pdf

S2 — Off-The-Shelf Software Use in Medical Devices
August 2023; final guidance  Official source
Local copy: fda/Off-The-Shelf-Guidance_0.pdf

S3 — Cybersecurity in Medical Devices: Quality Management System Considerations and Content of Premarket Submissions
February 2026; final guidance  Official source
Local copy: fda/FDA-Cybersecurity-QMS-Premarket-2026.pdf

S4 — Applying Human Factors and Usability Engineering to Medical Devices
August 2026; final guidance  Official source
Local copy: fda/FDA-Applying-Human-Factors-Usability-Engineering-2026.pdf

S5 — Content of Human Factors Information in Medical Device Marketing Submissions
May 2026; final guidance  Official source
Local copy: fda/FDA-Content-Human-Factors-Information-Marketing-Submissions.pdf

Local filenames identify copies in the sibling standards repository. Source publications remain there. Draft AI lifecycle guidance is background only. A process standard does not automatically require every internal record in a submission.

# Source pointers — 2

Use the sections cited with each topic. These are reading pointers for authored checks, not a clause-by-clause compliance checklist. Confirm the applicable edition, recognition and evidence route for the case.

S6 — Appropriate Use of Voluntary Consensus Standards in Premarket Submissions for Medical Devices
September 2018; final guidance  Official source
Local copy: fda/FDA-Appropriate-Use-Voluntary-Consensus-Standards-2018.pdf

S7 — FDA Recognized Consensus Standards Database
Live database; verify for the review date  Official source
Local copy: fda/FDA-Recognized-Consensus-Standards-Database-page.html

S8 — Marketing Submission Recommendations for a Predetermined Change Control Plan for Artificial Intelligence-Enabled Device Software Functions
August 2025; final guidance  Official source
Local copy: fda/Guidance-Predetermined-Change-Control-AI.pdf

S9 — Artificial Intelligence-Enabled Device Software Functions: Lifecycle Management and Marketing Submission Recommendations
January 2025; DRAFT, not for implementation  Official source
Local copy: fda/FDA-AI-Enabled-Device-Software-Functions-2025.pdf

S10 — IEC 62304 — Medical device software: Software life cycle processes
2006 + Amendment 1:2015; consolidated edition 1.1, final text
Local copy: iec62304{ed1.1}b.pdf

Local filenames identify copies in the sibling standards repository. Source publications remain there. Draft AI lifecycle guidance is background only. A process standard does not automatically require every internal record in a submission.

# Source pointers — 3

Use the sections cited with each topic. These are reading pointers for authored checks, not a clause-by-clause compliance checklist. Confirm the applicable edition, recognition and evidence route for the case.

S11 — ISO 14971 — Medical devices: Application of risk management to medical devices
2019; third edition
Local copy: ISO+14971-2019.pdf

S12 — IEC 62366-1 — Application of usability engineering to medical devices
2015 + Amendment 1:2020; consolidated edition 1.1, final text
Local copy: iec62366-1{ed1.1}b.pdf

S13 — IEC 81001-5-1 — Security: Activities in the product life cycle
2021; edition 1.0; read with interpretation sheet S17
Local copy: IEC+81001-5-1-2021.pdf

S14 — Premarket Approval Application and Humanitarian Device Exemption Modular Review
January 2025; final guidance  Official source
Local copy: fda/FDA-PMA-Modular-Review.pdf

S15 — Multiple Function Device Products: Policy and Considerations
July 2020; final guidance; local copy includes a QMSR notice  Official source
Local copy: fda/FDA-Multiple-Function-Device-Products-2020.pdf

Local filenames identify copies in the sibling standards repository. Source publications remain there. Draft AI lifecycle guidance is background only. A process standard does not automatically require every internal record in a submission.

# Source pointers — 4

Use the sections cited with each topic. These are reading pointers for authored checks, not a clause-by-clause compliance checklist. Confirm the applicable edition, recognition and evidence route for the case.

S16 — Postmarket Management of Cybersecurity in Medical Devices
December 2016; final guidance; local copy includes a QMSR notice  Official source
Local copy: fda/FDA-Postmarket-Management-Cybersecurity-2016.pdf

S17 — IEC 81001-5-1:2021 — Interpretation Sheet 1
November 2025; software categories, risk transfer and security maintenance
Local copy: IEC+81001-5-1-2021+ISH1-2025.pdf

Local filenames identify copies in the sibling standards repository. Source publications remain there. Draft AI lifecycle guidance is background only. A process standard does not automatically require every internal record in a submission.